Capilia® TB Testing Flow chart

Preparation for Capilia

Use a positive culture from MGIT broth or LJ slant or 7H11/7H10 media, confirm that the culture is AFB positive by ZN staining

For solid media dispense 0.2mL of Capilia extraction buffer into a sterile 2 mL centrifuge or cryovial tube

For solid media, use a sterile disposable loop to remove a loop full of culture material from the solid media

Transfer all the culture material from the loop to the extraction buffer aliquot

Ensure lid is secured, vortex suspension from solid media vigorously for 30 seconds

For MGIT cultures, insure cap is secure and votex entire culture vigorously for 30 seconds

Use mycobacterial suspension or MGIT culture to perform Capilia

Capilia Test Procedure

Remove the test cartridge from the foil wrapper and label with specimen identifications number

Deliver 100μL of the mycobacterial suspension prepare above to the culture placing area of the test plate

Set timer for 15 minutes; observe the reading area of the test cartridge after 15 minutes, but no more than 60 minutes and interpret the result

Positive test result:

Purple/red line is observed

On the reading areas of both [T] and [C] after 15 minutes, but no more than 60 minutes

Negative test Result:

Purple/red line is not observed on the reading area [T] but the color is observed on the reading area [C] after 15 minutes, but no more than 60 minutes

Invalid Test Result:

If no line is observed in the reading area [C] within 60 minutes, retest using a different plate