## Slide 1
A. Day of Positivity = Day 0
Incubate at least one more day after positivity.
B. If tube is ‘Day 1’ or ‘Day 2’, vortex 2-3 minutes, let settle for 20-30 minutes. Proceed to Preparation & Inoculation.
C. If tube is ‘Day 3’ to ‘Day 5’, vortex 2-3 minutes, let settle for 20-30 minutes. Dilute 1 mL of positive broth into 4 mL sterile saline. (1:5 dilution)
From a Positive 7 mL MGIT Tube:
From Solid Media:
Day 3 to 5
1:5 dilution
A
B
C
Day 0
Incubate for one day
Day 1 or 2
No dilution required
SIRE Inoculum
Sterile Tube with 4 mL MGIT broth and sterile glass beads
Solid Medium with growth < 14 days old
Prepare suspension
> 1.0 McFarland
Vortex 2 - 3 minutes
20 minutes for large clumps to settle
Transfer supernatent
15 minutes for clumps to settle
Transfer supernatent
Adjust to 0.5 McFarland turbidity
1 mL
4 mL sterile saline
1:5 dilution
SIRE Inoculum
1) Use growth on solid medium < 14 days old.
2) Make suspension >1.0 McFarland turbidity standard in MGIT broth containing 8 - 10 sterile glass beads.
3) Vortex 2 - 3 minutes.
4) Allow to sit 20 minutes for large clumps to settle.
5) Transfer supernatent fluid to another sterile tube and allow to sit for another 15 minutes.
6) Transfer supernatent fluid to another sterile tube and adjust turbidity to 0.5 McFarland turbidity standard.
7) Dilute 1 mL of final suspension in 4 mL sterile saline – inoculum tube. Vortex.
ILB500G09B EffectiveDate:05 June 2009
MGIT™ 960 SIRE Preparation & Inoculation

## Slide 2
1. Add 0.8 mL BACTEC™ MGIT™ 960 SIRE Supplement to each tube.
Growth
Control
#1 #2 #3 #4
0.8 mL
SIRE Supplement
3. Prepare Growth Control Inoculum: Pipette 0.1 mL of organism suspension into 10 mL sterile saline to make a 1:100 dilution. Vortex.
Sterile Saline
0.1 mL organism suspension
2. Add 100 µL of each drug to the appropriately labeled MGIT tubes (Strep to Strep tube and so on)
Growth
Control
#1 #2 #3 #4
S I R E
100 µl Drug
4. Inoculate Growth Control tube :Remove 0.5 mL of inoculum of the1:100 dilution. Invert to mix.
5. Inoculate drug containing tubes: Remove 0.5 mL of inoculum from the top 1/3 of the organism suspension (undiluted or 1:5 dilution). Invert to mix.
6. Place tubes into appropriate AST Set Carrier and enter into BACTEC MGIT 960 instrument.
Growth
Control
#1 #2 #3 #4
S I R E
0.5 mL of 1:100 dilution
0.5 mL of Organism Suspension
MGIT™ 960 SIRE Preparation & Inoculation