## Slide 1: BACTEC MGIT 960 DST quality control
<number>

## Slide 2: Why QC reagents?
<number>
BD thoroughly tests all BACTEC MGIT 960 DST products prior to release
So why test at your facility?
  Product may be mishandled during shipment
    Antibiotic potency may decrease
  Checks ability of laboratory to perform test
    Reagent storage
    Reagent preparation
    Organism preparation

## Slide 3: BACTEC™ MGIT™ 960 DST User quality control
<number>
New shipment or lot number of drug kits before use with patient isolates
Weekly SIRE and PZA DST QC
  If validated, otherwise set up a QC DST set each time DSTs are inoculated
QC organism:
  M. tuberculosis ATCC 27294 (H37Rv)
  SIRE drugs
  PZA
Drug concentrations tested should be the same as those used for routine testing
  Two concentrations of INH used for patients; QC both concentrations

## Slide 4: QC organisms
<number>
Pure culture of M. tuberculosis ATCC 27294 (H37RV) or well-characterized pan-susceptible MTB strain
  Colonies from solid media  14 days old
  MGIT 960 tube 1–5 days after flagged positive by instrument
Additional organisms may be tested to supplement BD QC recommendations
  Laboratory dependent
  May want to inoculate a known mono-resistant strain of MTB

## Slide 5: Drug and organism preparation
<number>
Care should be taken to ensure proper reconstitution of lyophilized drugs
  Wrong drug concentration
  Improper preparation significantly impacts results
Proper dilution of organism for drug and growth control tube is critical
  Suspension must be well mixed and homogeneous without clumps
  Prepare QC organism suspension in the same way as patient isolate suspension

## Slide 6: Inoculate SIRE and PZA DSTs
Drugs
  Rehydrate drugs with sterile water
  Remove aliquots of prepared drugs from freezer
  Check expiration dates
Prepare inoculum and dilutions
  Solid seeds
  Liquid seeds
<number>

## Slide 7: SIRE and PZA QC
<number>
Entry of tubes into MGIT 960
  Scan DST set carrier into the BACTEC MGIT 960 instrument
Interpretation of results
  Instrument will interpret results between days 4–13
  Results should read susceptible for all drugs
  If proper results are not observed, repeat test

## Slide 8: SIRE and PZA QC records
<number>
Record lot numbers of MGIT 960 tubes, drugs and drug supplements
Record QC results
Maintain records for a minimum of two years

## Slide 9: QC summary
<number>
QC of DST is critical to ensure test is functioning properly
  Should be performed on all new lots or new shipment of SIRE and PZA drugs
QC organism suspension can be prepared from MGIT tube or solid media
Organism suspension must be homogeneous, i.e. without clumps
Sterile tubes and pipettes must be used
Use calibrated pipettes to inoculate drugs to media tubes, not disposable transfer pipettes

## Slide 10: Acknowledgement
<number>