## Slide 1: Inoculum and drug preparation for the MGIT 960 DST system
Notes: <number>

## Slide 2: Liquid-based DST systems
Liquid methods incorporate a means of quantifying growth in the presence of a drug compared to a growth control with no drug
Growth characteristics in the medium are used as the basis for evaluation (drug vs no drug)
This growth characteristic is observed for the entire organism population and for 1% of the organism population
Notes: <number>
… read slide …

## Slide 3: Principle of MGIT DST
Isolated cultures from TB patients are subjected to growth in the presence of known concentration of a test drug

## Slide 4: SIRE susceptibility testing: MGIT 960
4–13 days qualitative test (from solid or liquid)
Increase in fluorescence interpreted as “Growth Units” (GU)
Susceptibility results and GU values are printed on the Unloaded AST Set report
The Instrument Inventory Report prints GU values for all ongoing AST sets
Notes: <number>

## Slide 5: Storage
Refrigerate lyophilized drugs at 2–8 oC
Reconstitute prior to use
Reconstituted drug solutions may be frozen in aliquots at -20 oC or lower and stored for up to 6 months or up to the date of original expiry, whichever comes sooner
Single thawing

## Slide 6: Reconstitution of SIRE lyophilized drugs
Reconstitute each critical concentration drug vial with 4 ml of sterile distilled/deionized water
Mix thoroughly and make sure drug is completely dissolved

## Slide 7: SIRE drug critical concentration
Drug Name | Critical Concentration
Streptomycin (STR) | 1.0 g/ml of medium
Isoniazid (INH) | 0.1 g/ml of medium
Rifampin (RIF) | 1.0 g/ml of medium
Ethambutol (EMB) | 5.0 g/ml of medium

## Slide 8: BACTEC MGIT 960 SIRE kit
Contents:
  8 vials of MGIT 960 SIRE supplement
    Do NOT exchange with MGIT growth supplement
  Critical concentrations of S, I, R, E drug in lyophilized form
No. of tests/kit = maximum of 100 tests

## Slide 9: Organism inoculum for MGIT 960 DST
MGIT 960 seeds
  Growth detection tubes from culture
  Growth detection subculture tubes
  Age of seeds from day of positivity
    Day 1 to day 5
Solid media
  Preparation of 0.5 McFarland organism
<number>

## Slide 10: SIRE inoculum preparation
Day 3 to 5
1:5 dilution
A
B
C
Day 0
Incubate for one day
Day 1 or 2
No dilution required
SIRE Inoculum
A. Day of positivity = Day 0
Incubate at least one more day after positivity
Day of positivity = 1–2,
Go directly to Preparation & Inoculation
C. Day of positivity = 3–5,
Dilute 1 ml of positive broth into 4 ml sterile saline
(1:5 dilution)

## Slide 11: SIRE inoculum from solid media
Colonies from solid media must be 14 days old from time of visible growth
Notes: <number>
… read slide …

## Slide 12: SIRE growth control suspension
Growth control:
  Make 1:100 dilution of working suspensions
  Inoculate 0.5 ml of 1:100 dilution to tube containing no drug
In a test tube, mix by inversion
0.1 ml culture suspension and
10.0 ml sterile saline

## Slide 13: Inoculation of SIRE
GC
S I R E
Prepare 5 MGIT tubes
Label MGIT tube
Add 0.8 ml SIRE supplement
Add 0.1 ml of corresponding drugs
Add 0.5 ml inoculum to drug tubes
Add 0.5 ml 1:100 dilution GC suspension to GC tube
Mix 5–6 times by inversion
Place in proper DST carriers and
into MGIT 960 instrument

## Slide 14: SIRE inoculation overview
Notes: <number>

## Slide 15: PZA susceptibility testing: MGIT 960
4–21 days qualitative test (from solid or liquid)
Increase in fluorescence interpreted as “Growth Units” (GU)
Susceptibility results and GU values are printed on the Unloaded DST Set report
The Instrument Inventory Report prints GU values for all ongoing DST sets
Notes: <number>

## Slide 16: PZA susceptibility testing
PZA is active only at low pH (in vitro)
  DST test must be performed in low pH environment
Solid media for PZA testing has pH of 5.0–5.5
  At this pH, many strains grow poorly or not at all
BACTEC PZA medium is around pH 5.9
PZA supplement required to enhance ability of organisms to grown at this pH
CLSI, M24-A
BACTEC PZA package insert
Notes: <number>

## Slide 17: PZA susceptibility test: MGIT 960
PZA medium is a modified 7H9 broth with a reduced pH of 5.9.
Supplement is enriched
  Different from SIRE AST supplement
Test at drug concentration 100 µg/ml
Results interpretable after 4–21 days
Qualitative test
Compared with BACTEC 460 Gold Standard (CLSI recommendations)
Notes: <number>

## Slide 18: PZA drug
Drug name | Approx. lyophilized drug/vial
Pyrazinamide (PZA) | 100 g

## Slide 19: Storage
Refrigerate lyophilized drugs at 2–8 oC
Reconstitute prior to use
Avoid exposure to light
Reconstituted drug solutions may be frozen in aliquots at -20 oC or lower and stored for up to 6 months or up to the date of original expiry, whichever comes sooner.
Single thawing

## Slide 20: Reconstitution of lyophilized drugs
Reconstitute PZA drug vial with 2.5 ml of sterile distilled/deionized water
Mix thoroughly and make sure drug is completely dissolved

## Slide 21: PZA inoculum preparation
A. Day of positivity = Day 0
Incubate at least one more day after positivity
Day of positivity = 1–2
Go directly to preparation & inoculation
C. Day of positivity = 3–5
Dilute 1 ml of positive broth into 4 ml sterile saline
(1:5 dilution)
Days 3 - 5
1:5 dilution
A
B
C
Day 0
Incubate for one day
Day 1 or 2
No dilution required
PZA Inoculum

## Slide 22: Inoculum from solid medium
1) Use growth on solid medium < 14 days old
2) Make suspension >1.0 McFarland turbidity standard in MGIT broth containing 8–10 glass beads
3) Vortex for 2–3 minutes
4) Allow to sit 20 minutes for large clumps to settle
5) Transfer supernatant fluid to another sterile tube and allow to sit for another 15 minutes
6) Transfer supernatant fluid to another sterile tube and adjust turbidity to 0.5 McFarland turbidity standard
7) Dilute 1 ml of final suspension in 4 ml sterile saline and mix.

## Slide 23: PZA growth control suspension
Growth control:
  Make 1:10 dilution of working suspensions
  Inoculate 0.5 ml of 1:10 dilution to tube containing no drug
In a test tube, mix by inversion
0.5 ml culture suspension
4.5 ml sterile saline

## Slide 24: Loading DST sets
Carry DST sets to the instrument in the DST transport rack
Open desired drawer and scan DST Set Carrier’s barcode label and isolate label on first tube
  Check DST set definition on instrument display and change if appropriate
Carefully insert the tube set into the indicated stations
  GC in the left-most position
Ensure that all the tubes AND carriers are fully seated in the drawer
Close drawer

## Slide 25: Summary points
Properly label tubes
Always add the exact volume of SIRE supplement, drugs and inoculum
Rehydrate drugs with the correct diluent and volume
Never add any drug to the GC
Verify pure culture
Check for mixed culture
Use cultures not older than 5 (for MGIT) and 14 (for solid) days
Avoid clumps, mix suspensions and inoculated vials properly
Notes: <number>

## Slide 26: Acknowledgement
<number>