# BACTEC MGIT 960 DST

OBJECTIVES: Upon completion of the DST module, the user will know how to:

- Prepare reagents and isolates for the inoculation of DST sets.

- Inoculate DST sets.

- Set the default DST Set definition.

- Enter DST Sets into the instrument.

- Remove completed DST Sets from the instrument.

- Resolve error/anonymous sets within the instrument.

- Print reports

- Interpret DST results.

## REAGENT PREPARATION

| NOTE: On receipt, store the lyophilized drug vials at 2 – 8°C. Once reconstituted, the antibiotic solutions may be frozen and stored at –20°C or colder for up to six months – not to exceed the original expiration date. Once thawed, use immediately. Discard any unused portions. |  |
| --- | --- |
| Materials: | - BACTEC MGIT 960 SIRE Kit containing one each lyophilized vials of Streptomycin, Isoniazid, Rifampin and Ethambutol, and eight vials of SIRE Supplement (approximately 40 tests per drug per kit). - BACTEC MGIT 960 STR 4.0 Kit containing one vial lyophilized Streptomycin and two vials of SIRE Supplement (approximately 20 tests per kit). - BACTEC MGIT 960 INH 0.4 Kit containing one vial lyophilized Isoniazid and two vials of SIRE Supplement (approximately 20 tests per kit). - BACTEC MGIT 960 EMB 7.5 Kit containing one vial lyophilized Ethambutol and two vials of SIRE Supplement (approximately 20 tests per kit). - BACTEC MGIT 960 PZA Kit containing two vials of Pyrazinamide and six vials of PZA Supplement (approximately 50 tests per kit) - BACTEC MGIT 960 PZA Medium 1 box containing 25 tubes. |
| Additional Materials: | - BBL™MGIT™ 7 mL Mycobacteria Growth Indicator Tube. - Ancillary culture media. - Reagents. - Quality control organisms. - Laboratory equipment as required for procedure. |

| # TASK | # INSTRUCTIONS |
| --- | --- |
| # Prepare BACTEC™ MGIT™ 960 SIRE Kit reagents. | - Reconstitute each BACTEC™ MGIT™ 960 SIRE Kit Streptomycin lyophilized drug vial with 4 mL of sterile distilled/deionized water to make a stock solution of 83 g/mL. - Reconstitute each BACTEC™ MGIT™ 960 SIRE Kit Isoniazid lyophilized drug vial with 4 mL of sterile distilled/deionized water to make a stock solution of 8.3 g/mL. - Reconstitute each BACTEC™ MGIT™ 960 SIRE Kit Rifampin lyophilized drug vial with 4 mL of sterile distilled/deionized water to make a stock solution of 83 g/mL. - Reconstitute each BACTEC™ MGIT™ 960 SIRE Kit Ethambutol lyophilized drug vial with 4 mL of sterile distilled/deionized water to make a stock solution of 415 g/mL. |
| # Prepare BACTEC™ MGIT™ 960 PZA KIT Reagents | - Reconstitute each BACTEC™ MGIT™ PZA kit lyophilized drug vial with 2.5 mL of sterile distilled/deionized water to make a stock solution of 8000 g/mL. |
| NOTE: The following are reconstituted with a different volume. Failure to use the appropriate volume of sterile distilled/deionized water for reconstitution of the higher drug concentration will invalidate these tests. |  |
| BACTEC™ MGIT™ 960 STR 4.0 Kit Reagent Preparation | - Reconstitute each BACTEC™ MGIT™ 960 STR 4.0 Kit Streptomycin lyophilized drug vial with 2 mL of sterile distilled/deionized water to make a stock solution of 332 g/mL. |
| BACTEC™ MGIT™ 960 INH 0.4 Kit Reagent Preparation | - Reconstitute each BACTEC™ MGIT™ 960 INH 0.4 Kit Isoniazid lyophilized drug vial with 2 mL of sterile distilled/deionized water to make a stock solution of 33.2 g/mL. |
| BACTEC™ MGIT™ 960 EMB 7.5 Kit Reagent Preparation | - Reconstitute each BACTEC™ MGIT™ 960 EMB 7.5 Kit Ethambutol lyophilized drug vial with 2 mL of sterile distilled/deionized water to make a stock solution of 622.5 g/mL. |

## SPECIMEN PREPARATION

POTENTIALLY INFECTIOUS TEST SPECIMEN: Observe “Universal Precautions” and institutional guidelines when handling and disposing of infectious material.

Working with M. tuberculosis growth in culture requires Biosafety Level (BSL) 3 practices, containment equipment and facilities.

| TASK | # INSTRUCTIONS |
| --- | --- |
| Preparation from a Positive 7 mL BACTEC MGIT Tube | NOTE: The MGIT tube should be used the day after it is instrument positive (Day 1) up to and including the 5th day (Day 5) after instrument positivity. - If tube is a Day 1 or Day 2 positive tube, then proceed to “DST Set Preparation and Inoculation” for the appropriate antimicrobic. - For Day 3-5 positive tubes: - Dilute 1 mL of positive broth in 4 mL of sterile saline (1:5 dilution) - Proceed to “DST Set Preparation and Inoculation” for the appropriate antimicrobic. |
| Preparation of the isolate from solid media. | - Confirm, by appropriate identification techniques, that the isolate to be tested is a pure culture. - Add 4 mL of Middlebrook 7H9 Broth (or MGIT broth) to a sterile tube with cap containing 8-10 glass beads. - With a sterile loop, scrape as many colonies as possible from growth up to 14 days old. Suspend colonies in the broth prepared above for a suspension > 1.0 McFarland standard. Note: be careful not to scrape up agar from the culture when removing colonies. - Vortex the suspension for 2-3 minutes to break up the larger clumps. - Let the suspension sit for 20 minutes without disturbing. - Transfer the supernatant fluid to another sterile tube with cap (avoid transferring any of the sediment) and let sit for another 15 minutes. - Transfer the supernatant fluid (it should be smooth, free of any clumps) to a third sterile tube. - Adjust suspension to a 0.5 McFarland standard by a visual comparison to a 0.5 McFarland turbidity standard. Note: Tube with supernatant and 0.5 McFarland standard tube should be of the same diameter to make accurate comparison. - Dilute 1 mL of the adjusted suspension in 4 mL of sterile saline (1:5 dilution). Proceed to “DST Set Preparation and Inoculation” for the appropriate antimicrobic |
| NOTE: Failure to manipulate and dilute the specimen in the correct manner will cause inaccurate results. Do not deviate from written set up procedure. |  |

## DST SET PREPARATION AND INOCULATION

| # TASK | # INSTRUCTIONS |
| --- | --- |
| DST Set Preparation for SIRE  Prior to use, the user should examine the tubes for evidence of contamination or damage. Discard any tubes which appear unsuitable. | - Label the appropriate number of 7 mL MGIT tubes – one for (Growth Control) GC and one for each drug being tested. Place patient accession on the GC tube. NOTE: SIRE MGIT tubes, reagents, and/or dilutions differ from PZA MGIT tubes, reagents, and/or dilutions. Failure to use the correct tube and/or reagents will give inaccurate results. - Place tubes in the correct sequence in the appropriate size DST Set Carrier for the DST Set you are testing. NOTE: Using a carrier of the incorrect size will cause the system to interpret the absence of a tube as an error condition, which invalidates the results for that DST Set. - Place DST Set Carrier in the DST Transport rack. (Each DST tube rack holds up to eight 5-tube DST sets along with three 8-tube sets.) It is strongly recommended that you prepare the DST specimens with the tubes in the tube rack. If a cap is not secure, and the DST Set Carrier is lifted, the affected tube will be captured by the tube rack. - Aseptically add 0.8 mL of BACTEC MGIT 960 SIRE Supplement to each tube. NOTE: SIRE supplement is different than PZA DST Supplement. If the incorrect supplement and/or tube is used, the results will be inaccurate. - Aseptically pipette 100 L of each drug into the appropriately labeled MGIT drug tube. NOTE: Failure to inoculate the correct drug to the assigned tube will result in accurate results. NOTE: No drug should be added to the Growth Control (GC) tube. |
| Growth Control Preparation for SIRE | - Aseptically pipette 0.1 mL of the organism suspension into 10 mL of sterile saline to prepare the 1:100 Growth Control suspension. - Mix thoroughly. - Inoculate 0.5 mL of the 1:100 Growth Control suspension into the “GC” tube. NOTE: Failure to properly dilute the growth control will lead to inaccurate results. |
| DST SIRE Set Inoculation | - Aseptically pipette 0.5 mL of the organism suspension into each drug containing tube. - Tightly recap the tubes and mix well. Make sure that all caps are screwed on the tubes securely before replacing them in the DST Set carrier. - Streak 0.1 mL of the organism suspension to a Sheep Blood Agar plate. Proceed to “DST Set Entry” |
| DST Set Preparation for PZA  Prior to use, the user should examine the tubes for evidence of contamination or damage. Discard any tubes which appear unsuitable. | - Label two 7 mL MGIT PZA tubes for each test isolate. NOTE: PZA tubes, reagents, and/or dilutions differ from SIRE tubes, reagents, and/or dilutions. Failure to use the correct tube and/or reagents will give inaccurate results. - Place tubes in the correct sequence in the two tube DST Set Carrier. NOTE: Using a carrier of the incorrect size will cause the system to interpret the absence of a tube as an error condition, which invalidates the results for that DST Set. - Place DST Set Carrier in the DST Transport rack. - (Each DST tube rack holds up to eight 5-tube DST sets along with three 8-tube sets.) It is strongly recommended that you prepare the DST specimens with the tubes in the tube rack. If a cap is not secure, and the DST Set Carrier is lifted, the affected tube will be captured by the tube rack. - Aseptically add 0.8 mL of BACTEC MGIT PZA Supplement to each tube. NOTE: PZA supplement is different than SIRE DST Supplement. If the incorrect supplement and/or tube is used, the results will be inaccurate. - Aseptically pipette 100 L of reconstituted 8000 g/mL BACTEC MGIT PZA drug solution to the appropriately labeled BACTEC MGIT PZA Tube. No drug solution should be added to the GC tube. NOTE: Use of an incorrect drug will lead to inaccurate results. |
| Growth Control Preparation for PZA | - Aseptically pipette 0.5 mL of the organism suspension into 4.5 mL of sterile saline to prepare the 1:10 Growth Control suspension. - Mix thoroughly. - Inoculate 0.5 mL of the 1:10 Growth Control suspension into the “GC” tube. NOTE: Failure to properly dilute the growth control will lead to inaccurate results. |
| PZA Set Inoculation | - Aseptically pipette 0.5 mL of the organism suspension into each drug containing tube. - Tightly recap the tubes and mix well. Make sure that all caps are screwed on the tubes securely before replacing them in the DST Set carrier. - Streak 0.1 mL of the organism suspension to a Sheep Blood Agar plate and incubate at 35 - 37C. Proceed to “DST Set Entry” |

## DST SET ENTRY

*** NOTE: DST functionality must be enabled and disabled by a Becton Dickinson representative

| # TASK | # INSTRUCTIONS |
| --- | --- |
| Set Default DST Definition Default set indicated by check mark | - Ensure that all instrument drawers are closed. - Press “configuration” soft key to enter configuration mode. - Press “configuration” soft key to access the “Set Default DST Set” function. - Press the “perform action” soft key. This will bring up a new series of soft keys that enable the user to select the 2-tube carrier, 3-tube carrier, etc. |
|  | Default DST sets can be configured for each DST Set Carrier size used (2 tubes, 3 tubes, etc. See BACTEC MGIT 960 User’s Manual for available DST Set Carrier definitions.) - Press the soft key corresponding to the desired carrier size. - The current default DST Set configuration, which is always set number 1 (of x), is shown. - Press the UP ARROW▲ or DOWN ARROW ▼ to scroll through the available DST Set definitions. - Select desired default set definition |
| Enter DST Sets | - Take DST tube sets to the instrument using the DST Transport Rack. - Open desired drawer. - Press the “tube entry” soft key. |
| # Enter DST Sets, (cont.) | - Scan the DST Set Carrier’s barcode label. The carrier barcode indicates that what you are entering is a DST Set, how many tubes are in the set, and the set’s sequence number. 43 95 00002041 - FIRST two digits: 92 = 2-tube carrier 43 = signifies a valid bd barcode - NEXT two digits: 92 = 2-tube carrier 93 = 3-tube carrier 94 = 4-tube carrier 95 = 5-tube carrier 96 = 6-tube carrier 97 = 7-tube carrier 98 = 8-tube carrier - remaining digits are unique to that specific DST Set Carrier - Check that the DST set definition and isolate number are displayed correctly for the DST Set currently being entered. The default DST Set is automatically displayed when a new DST carrier barcode is scanned. - If accession barcode is enabled, scan accession barcode for the DST Set. If accession barcode is not available, press the accession barcode not available” soft key. - If set definition is incorrect, go to “Change DST Set Definition”. NOTE: Use of an incorrect DST set definition will cause inaccurate reports/results. - If isolate number is incorrect, press the “isolate number” soft key and press UP ARROW ▲ key to increase or DOWN ARROW ▼ key to decrease the isolate number. - Carefully insert the DST Set in the stations indicated by GREEN indicators, with the Growth Control tube leftmost of the tubes in the set. NOTE: Be sure to fully seat all tubes in the correct order. |
| NOTE: Make sure tubes are placed in the carrier in the exact order indicated on the LCD display. This is necessary for accurate results to be reported at the DST test completion. |  |
| Change DST Set Definition | - After scanning the DST carrier and the accession number (if available), if the DST set definition is incorrect, go to step 2. - Press the “change DST set definition” soft key. *Select correct carrier size. - Use the UP ARROW ▲ key or DOWN ARROW ▼ key to scroll through the available definitions for the scanned carrier size/drug order. The default set is indicated by a check mark in the main body of the display. |

## DST SET REMOVAL

| # TASK | # INSTRUCTIONS |
| --- | --- |
| Remove completed DST Sets | - Open desired drawer. - Press the “remove completed DST sets” soft key. - The first completed DST Set stations illuminate with FLASHING GREEN indicators. - Remove the carrier and scan its barcode label. - The LEDs at this station extinguish. - Repeat to remove additional completed DST Sets. - Place completed DST sets in the DST Transport Rack. |